Inspection Report

1 June 2026

Healthy Smiles Independent Dental Hygiene Clinic

Type of service: Independent Hospital (IH) – Dental Treatment

Address: 63-65 Castle Street, Ballycastle, BT54 6AS

Telephone number: 075 1415 6986

Information on legislation and standards underpinning inspections can be found on our website https://www.rqia.org.uk/, The Independent Health Care Regulations (Northern Ireland) 2005 and Minimum Standards for Dental Care and Treatment (March 2011).

1.0	Service information

| Applicant Organisation: Healthy Smiles Independent Dental Hygiene Clinic Ltd Applicant Responsible Person: Mrs Shauna McHenry | Applicant Registered Manager: Mrs Shauna McHenry Date registered: Awaiting registration |
| --- | --- |
| Person in charge at the time of inspection: Mrs Shauna McHenry | Number of registered places: One |
| Categories of care: Independent Hospital (IH) – Dental Treatment |  |
| Brief description of how the service operates: Healthy Smiles Independent Dental Hygiene Clinic has made an application to register with the Regulation and Quality Improvement Authority (RQIA) as an independent hospital (IH) with a dental treatment category of care. The application was for the registration of one dental chair providing private dental hygienist care and treatment without conscious sedation. |  |

2.0	Inspection summary

This was an announced pre-registration inspection, undertaken by a care inspector on 1 June 2026 from 10.00 am to 1.00 pm.

A desktop review of the premises, medicines management and finance components of the registration application was also undertaken by RQIA estates, pharmacy and finance inspectors respectively.  All have confirmed approval of the registration application from their corresponding estates, pharmacy and finance perspectives.

The inspection sought to assess an application submitted to RQIA for the registration of Healthy Smiles Independent Dental Hygiene Clinic as an independent hospital providing dental treatment.  The application was for the registration of one dental chairs.

An application was also submitted for the registration of Mrs Shauna McHenry as the responsible individual and as the registered manager.

There was evidence of good practice in relation to the recruitment and selection of staff; staff training; management of medical emergencies; infection prevention and control (IPC); decontamination of reusable dental instruments; radiology and radiation safety; management of complaints and incidents; and governance arrangements.

No immediate concerns were identified regarding the delivery of front line patient care.

Information has been gathered throughout the registration process.  Scrutiny of this information means that registration of this dental practice is approved.

Mrs McHenry has agreed to the following condition being placed on the registration of Healthy Smiles Independent Dental Hygiene Clinic:

“Mrs McHenry agrees to inform RQIA when the “Patient Group Direction (PGD) for the Administration of Salbutamol in Dental Emergencies” is in place.  Until the PGD is in place, all patients with asthma will be requested to bring their salbutamol inhaler to their appointment.”

Information in relation to this is included in Section 5.2.6 of this report.

A certificate of registration will be issued to Healthy Smiles Independent Dental Hygiene Clinic following the approval of the registration application.  Mrs McHenry is aware that the RQIA certificate of registration must be displayed in a prominent place.

3.0	How we inspect

RQIA is required to undertake pre-registration inspections following receipt of service applications.  To do this, we review the registration application and supporting documents, meet and talk with staff and management and observe practice on the day of the inspection.

The information obtained is then considered before a determination is made on whether the practice is operating in accordance with the relevant legislation and minimum standards and can be approved for registration.

Before the inspection a range of information relevant to the registration application was reviewed.  This included the following records:

- application to register the practice

- application to register the responsible individual

- application to register the manager

- the proposed statement of purpose

- the proposed patient guide

During the inspection a tour of the premises was undertaken and the inspector met with Mrs McHenry and a dental nurse.

A sample of records in relation to the following areas was also reviewed:

- staffing

- recruitment and selection

- safeguarding

- medical emergencies

- infection prevention and control (IPC) and decontamination

- radiology and radiation safety

- records management

- management and governance arrangements

- equipment and premises maintenance arrangements

Examples of good practice are acknowledged and any areas for improvement are discussed with the person in charge and detailed in the quality improvement plan (QIP).

4.0	What people told us about the care and treatment?

As this a new dental practice, not yet operational, it was not possible to consult with patients however; this area will be reviewed during future inspections.  Mrs McHenry confirmed that patient satisfaction surveys will be undertaken annually.  This area is further discussed in section 5.2.13.

5.0	The inspection

5.1	What action has been taken to meet any areas for improvement identified at or since last inspection?

As this is a pre-registration inspection there are no previous inspections for Healthy Smiles Independent Dental Hygiene Clinic.

5.2	Inspection findings

5.2.1	Has the statement of purpose been developed in keeping with Regulation 7 Schedule 1 of the regulations?

A review of the proposed statement of purpose identified that it fully reflected the key areas and themes specified in Regulation 7, Schedule 1 of The Independent Health Care Regulations (Northern Ireland) 2005.  Mrs McHenry is aware that the statement of purpose should be reviewed and updated as and when necessary.

5.2.2	Has the patient guide been developed in keeping with Regulation 8, of the regulations?

A review of the proposed patient guide identified that it fully reflected the key areas and themes specified in Regulation 8 of The Independent Health Care Regulations (Northern Ireland) 2005.  Mrs McHenry is aware that the patient guide should be reviewed and updated as and when necessary.

5.2.3 	How does this service ensure that staffing levels are safe to meet the needs of patients and staff are appropriately trained to fulfil the duties of their role?

Staffing arrangements were reviewed and it was confirmed that there are appropriately skilled and qualified staff involved in the delivery of services.

# Dental practices are required to maintain a staff register.  A staff register was in place which was noted to be up to date and included all the required information.  Mrs McHenry was aware that the staff register is a live document and should be updated and amended as and when required.

Discussion with Mrs McHenry and review of documentation identified that arrangements were in place for monitoring the General Dental Council (GDC) registration status of all clinical staff on appointment and on an ongoing basis.  Arrangements were also in place to review individual indemnity insurance of clinical members of the dental team, where appropriate, on an annual basis.  It was confirmed that a record of these checks are retained and available for inspection.

It was confirmed that all staff providing services in the practice, including visiting professionals, are either directly employed by the practice and have a contract of employment or have an associate agreement in place.

The dental team take part in ongoing training to update their knowledge and skills, relevant to their role.  Induction programmes relevant to roles and responsibilities are required to be completed when new staff join the team and it was confirmed that an induction programme would be provided to any new staff recruited.

# A system was in place to monitor all aspects of ongoing professional development and a record was retained of all training and professional development activities.  A review of the records confirmed that all current staff had undertaken training in keeping with RQIA training guidance and legislation.

It was demonstrated that staffing levels meet the needs of patients and staff are appropriately trained to fulfil the duties of their role.

5.2.4	Does the practice’s recruitment and selection procedure comply with relevant legislation?

# There were recruitment and selection policies and procedures in place.  A review of these identified that they required further development to fully reflect the legislation and best practice guidance.  This was discussed with Mrs McHenry and advice provided.  Mrs McHenry agreed to address this matter following the inspection.

# Mrs McHenry oversees the recruitment and selection of the dental team, and approves all staff appointments.  Discussion with Mrs McHenry confirmed that she had a clear understanding of the legislation and best practice guidance.

# The personnel file of a newly recruited staff member was reviewed and evidenced that all relevant records had been sought; reviewed and stored as required under Schedule 2 of The Independent Health Care Regulations (Northern Ireland) 2005, as amended.

There was evidence of job descriptions and induction checklists for the different staff roles.  A review of records confirmed that if a professional qualification is a requirement of the post, a registration check is made with the appropriate professional regulatory body.

# Discussion with a dental team member confirmed that they have been provided with a job description upon commencing work at the practice.

# It was demonstrated that robust arrangements were in place for the recruitment of dental team members to comply with legislation and best practice guidance.

5.2.5 	Does the dental team meet current best practice guidance for the management of safeguarding concerns?

# A review of the arrangements for safeguarding evidenced that policies and procedures were in place for the safeguarding and protection of adults and children at risk of harm.

# Review of the policies and procedures demonstrated that all relevant and required information was included.  These policies and procedures included the types and indicators of abuse and distinct referral pathways in the event of a safeguarding issue arising with an adult or child.  The relevant contact details for onward referral to the local Health and Social Care Trust (HSCT), should a safeguarding issue arise, were also included.

# Mrs McHenry demonstrated a good awareness of the types and indicators of abuse and the actions to be taken in the event of a safeguarding issue being identified, including who the nominated safeguarding lead was.

# A review of training records evidenced that the dental team had received training in safeguarding of adults and children, as outlined in The Minimum Standards for Dental Care and Treatment (2011).  Mrs McHenry was aware of the frequency of safeguarding update training.

# It was confirmed with Mrs McHenry, the safeguarding lead within the establishment, that she had completed formal level two training in safeguarding adults, in keeping with the Northern Ireland Adult Safeguarding Partnership training strategy (revised 2016) and the Safeguarding Board for Northern Ireland (SBNI) Child Safeguarding Learning and Development Strategy and Framework 2020 – 2023 (July 2021).

A copy of the regional policy entitled ‘Co-operating to Safeguard Children and Young People in Northern Ireland’ (November 2024) and a copy of the regional guidance document entitled ‘Adult Safeguarding Prevention and Protection in Partnership’ (July 2015) were available for reference.

It was demonstrated that appropriate arrangements were in place to manage a safeguarding issue should it arise.

5.2.6 	Is the practice fully equipped and are the dental team trained to manage medical emergencies?

The British National Formulary (BNF) and the Resuscitation Council (UK) specify the emergency medicines and medical emergency equipment that must be available to safely and effectively manage a medical emergency.

# There was a medical emergency policy and procedure in place and a review of this evidenced that it was comprehensive, reflected legislation and best practice guidance.  Protocols were available to guide the dental team on how to manage recognised medical emergencies.

Robust systems were in place to ensure that emergency medicines and equipment do not exceed their expiry date and are immediately available.

Managing medical emergencies is included in the dental team induction programme and training is updated annually.  The records reviewed verified that the staff had completed medical emergency refresher training on an annual basis.

Members of the dental team were able to describe the actions they would take, in the event of a medical emergency, and were familiar with the location of medical emergency medicines and equipment.

As discussed previously, an RQIA pharmacy inspector conducted a desktop review of the medicines management component of the application.

A review of documentation and discussion with Mrs McHenry identified that a PGD is required for the administration of Salbutamol.  Mrs McHenry has taken action to obtain the required PGD and will confirm with RQIA once this is in place.

Fulfilment of this condition will ensure that the management of medical emergencies complies with legislative requirements, professional standards and guidelines.

5.2.7	Does the dental team provide dental care and treatment using conscious sedation in line with the legislation and guidance?

Conscious sedation helps reduce anxiety, discomfort, and pain during certain procedures.  This is accomplished with medications or medical gases to relax the patient.

Mrs McHenry confirmed that conscious sedation is not offered in Healthy Smile Independent Dental Hygiene clinic.

5.2.8	Does the dental team adhere to infection prevention and control (IPC) best practice guidance?

The IPC arrangements were reviewed throughout the practice to evidence that the risk of infection transmission to patients, visitors and staff was minimised.

There was an overarching IPC policy and associated procedures in place.  A review of these documents demonstrated that they reflected legislation and best practice guidance, including guidance from the Health and Social Care Public Health Agency and the Infection Prevention and Control Manual for Northern Ireland.  Mrs McHenry confirmed arrangements are in place to check Department of Health (DoH) websites for further advisory information, guidance and alerts in this regard.

Mrs McHenry confirmed she was the nominated lead, who had responsibility for IPC and decontamination in the practice.  It was confirmed that Mrs McHenry had undertaken IPC and decontamination training in line with her continuing professional development and had retained the necessary training certificates as evidence.

During a tour of some areas of the practice, it was observed that clinical and decontamination areas were clean, tidy and uncluttered.  All areas of the practice observed were equipped to meet the needs of patients.

The arrangements for personal protective equipment (PPE) were reviewed and it was noted that appropriate PPE was readily available for the dental team in accordance with the treatments provided.

Using the Infection Prevention Society (IPS) audit tool, IPC audits are routinely undertaken by members of the dental team to self-assess compliance with best practice guidance.  The purpose of these audits is to assess compliance with key elements of IPC, relevant to dentistry, including the arrangements for environmental cleaning; the use of PPE; hand hygiene practice; and waste and sharps management.  This audit also includes the decontamination of reusable dental instruments which is discussed further in the following section of this report.  Mrs McHenry confirmed that this IPC audit will be completed prior to opening and thereafter on a six monthly basis.  Where applicable, an action plan will be generated to address any improvements required.

Hepatitis B vaccination is recommended for clinical members of the dental team as it protects them if exposed to this virus.  A system was in place to ensure that relevant members of the dental team have received this vaccination.  Mrs McHenry confirmed that vaccination history is checked during the recruitment process and vaccination records are retained in personnel files.

A review of training records evidenced that the dental team had completed relevant IPC training and had received regular updates.

Review of IPC arrangements evidenced that the dental team adheres to best practice guidance to minimise the risk of infection transmission to patients, visitors and staff.

5.2.9	Does the dental team meet current best practice guidance for the decontamination of reusable dental instruments?

Robust procedures and a dedicated decontamination room must be in place to minimise the risk of infection transmission to patients, visitors and staff in line with Health Technical Memorandum 01-05: Decontamination in primary care dental practices, (HTM 01-05), published by the DoH.

There was a range of policies and procedures in place for the decontamination of reusable dental instruments that were comprehensive and reflected legislation, minimum standards and best practice guidance.

There was a designated decontamination room separate from patient treatment areas and dedicated to the decontamination process.  The design and layout of this room complied with best practice guidance and the equipment was sufficient to meet the requirements of the practice.  Records evidencing that the equipment for cleaning and sterilising instruments was inspected, validated, maintained and used in line with the manufacturers’ guidance were reviewed.  Review of equipment logbooks demonstrated that all required tests to check the efficiency of the machines had been undertaken.

Discussion with Mrs McHenry and a dental nurse confirmed that they had received training on the decontamination of reusable dental instruments in keeping with their role and responsibilities.  They demonstrated good knowledge and understanding of the decontamination process and were able to describe the equipment treated as single use and the equipment suitable for decontamination.

Decontamination arrangements demonstrated that the dental team are adhering to current best practice guidance on the decontamination of dental instruments.

5.2.10 How does the dental team ensure that appropriate radiographs (x-rays) are taken safely?

The arrangements regarding radiology and radiation safety were reviewed to ensure that appropriate safeguards were in place to protect patients, visitors and staff from the ionising radiation produced by taking an x-ray.

Dental practices are required to notify and register any equipment producing ionising radiation with the Health and Safety Executive (HSE) (Northern Ireland).  A review of records evidenced the practice had registered with the HSE.

The practice has one surgery which has an intra-oral x-ray machine and the equipment inventory reflected this.

A radiation protection advisor (RPA), medical physics expert (MPE) and radiation protection supervisor (RPS) have been appointed in line with legislation.  A dedicated radiation protection file containing the relevant local rules, employer’s procedures and other additional information was retained.

A review of the file confirmed that the Employer had entitled the dental team to undertake specific roles and responsibilities associated with radiology and ensured that these staff had completed appropriate training.  The RPS oversees radiation safety within the practice and regularly reviews the radiation protection file to ensure that it is accurate and up to date.

The appointed RPA must undertake a critical examination and acceptance test of all new x-ray equipment.  A review of documentation evidenced that a critical examination and acceptance test (CEAT) report for the new intra-oral x-ray was undertaken on during March 2026 and the recommendations, made by the RPA, had been actioned.

A copy of the local rules was on display near the x-ray machine and staff had signed to confirm that they had read and understood these.

Quality assurance systems and processes were in place to ensure that all matters relating to x-rays reflect legislation and best practice guidance.  Discussion with Mrs McHenry evidenced that all measures would be taken to optimise radiation dose exposure.  This included the use of rectangular collimation, x-ray audits and digital x-ray processing.

Discussion with Mrs McHenry confirmed that arrangements were in place to service and maintain the x-ray equipment in accordance with manufacturer’s instructions.

The radiology and radiation safety arrangements evidenced that procedures are in place to ensure that appropriate x-rays are taken safely.

5.2.11 Are arrangements in place to maintain the environment?

A tour of the premises confirmed that the environment was maintained to a very good standard of maintenance and decor.

Detailed cleaning schedules were in place for all areas which were signed on completion.  A colour coded cleaning system was in place to minimise the risk of cross-contamination.

Suitable arrangements were in place for maintaining the environment.  Mrs McHenry confirmed, on the submitted self-assessment, that the relevant risk assessments were in place.

As discussed in Section 2.0 of this report, an RQIA estates inspector undertook a desktop review of the premises section of the registration application and approved the registration application from an estates perspective.

5.2.12	Are arrangements in place to ensure complaints and incidents will be effectively managed?

The arrangements for the management of complaints and incidents were reviewed to ensure that they will be managed in keeping with legislation and best practice guidance.

The complaints policy and procedure provided clear instructions for patients and staff to follow.  Patients and/or their representatives are made aware of how to make a complaint by way of the patient’s guide and information on display in the practice.

Arrangements were in place to record any complaint received in a complaints register and retain all relevant records including details of any investigation undertaken, all communication with complainants, the outcome of the complaint and the complainant’s level of satisfaction.

Discussion with Mrs McHenry confirmed that complaints will be managed in accordance with best practice guidance and that arrangements are in place to undertaken a complaints audit to identify trends, drive quality improvement and to enhance service provision.

# An incident policy and procedure was in place which includes the reporting arrangements to RQIA.

# Mrs McHenry confirmed that incidents will be effectively documented and investigated in line with legislation.  All relevant incidents will be reported to RQIA and other relevant organisations in accordance with legislation and RQIA’s Statutory Notification of Incidents and Deaths.  Adverse incidents will be audited to identify trends and improve service provided.

The dental team was knowledgeable on how to deal with and respond to complaints and incidents in accordance with legislation, minimum standards and the DoH guidance.

Arrangements were in place to share information with the dental team about complaints and incidents including any learning outcomes, and also compliments received.

Systems were in place to ensure that complaints and incidents will be managed effectively in accordance with legislation and best practice guidance.

5.2.13 Are arrangements in place to review the quality of treatment and other services?

The arrangements for the review of the quality of treatment and other services were reviewed to ensure that systems were in place in keeping with legislation and best practice guidance.

Mrs McHenry told us that patients will have the opportunity to provide feedback.  This feedback will be gathered in the form of either satisfaction surveys, social media channels or in the form of verbal feedback.

Advice was given to retain information regarding patient feedback in an anonymised format. This should be summarised and an annual report generated.  This report should be made available to patients and other interested parties.  Mrs McHenry agreed to implement this upon registration.

A number of quality assurance systems and processes will be implemented.  These include IPC and radiology audits, staff appraisals and review of ongoing professional development.

It was demonstrated that arrangements are in place to review the quality of treatment and other services in keeping with legislation and best practice guidance.

5.2.14	Are policies and procedures being effectively managed?

It was confirmed that policies and procedures are available for staff reference.  Policies and procedures are centrally indexed and are retained in a manner making them accessible to staff.

Mrs McHenry confirmed that policies and procedures are signed by herself as the author with the date of implementation and review included.  Policies and procedures will be subject to a systematic three yearly review or as changes occur.

A systematic approach had been taken and staff had signed to confirm that they had read a number of policies to date.

It was demonstrated that policies and procedures have been developed in accordance with legislation and best practice.

5.2.15	Are records being managed in accordance with legislation and best practice guidance?

A review of the arrangements in respect of records management demonstrated that electronic records are maintained.  These have different levels of access afforded to staff, dependent on their role and responsibilities.  Appropriate systems and processes were in place for the management of records and maintaining patient confidentiality.

Healthy Smiles Independent Dental Hygiene Clinic is registered with the Information Commissioner’s Office (ICO).

Mrs McHenry was able to describe the arrangements in place for effective records management and demonstrated an awareness of the need to comply with the General Data Protection Regulations (GDPR) 2018.

The management of records within the practice was found to be in line with legislation and best practice.

5.2.16	How does a registered provider who is not in day to day management of the practice assure themselves of the quality of the services provided?

Where the business entity operating a dental practice is a corporate body or partnership or an individual owner who is not in day to day management of the practice, unannounced quality monitoring visits by the registered provider must be undertaken and documented every six months; as required by Regulation 26 of The Independent Health Care Regulations (Northern Ireland) 2005.

Mrs McHenry will be in day to day management of the practice, therefore the unannounced quality monitoring visits by the registered provider are not applicable.

5.2.17	Does the dental team have suitable arrangements in place to record equality data?

The arrangements in place in relation to the equality of opportunity for patients and the importance of staff being aware of equality legislation and recognising and responding to the diverse needs of patients was discussed with Mrs McHenry.

5.2.18	Fit person interview

Providers of regulated establishments require to be registered with RQIA in accordance with Article 12 of The Health and Personal Social Services (Quality, Improvement and Regulation) (Northern Ireland) Order 2003, as it is an offence to carry on an establishment of any description without being registered in respect of it.

Mrs McHenry submitted to RQIA an application on behalf of Healthy Smiles Independent Dental Hygiene Clinic Ltd to become the responsible individual and registered manager of Healthy Smiles Independent Dental Hygiene Clinic.  The relevant information, supporting documentation and appropriate fees accompanied the application.

A fit person interview was undertaken as part of the pre-registration inspection.  Discussion with Mrs McHenry evidenced that she had a clear understanding of her role and responsibilities as a responsible individual and registered manager under the relevant legislation and minimum standards.  The following issues were discussed:

- the statement of purpose and patient guide

- management of complaints

- notification of untoward incidents to RQIA and other relevant bodies

- notification of registered person’s/manager absences, change of ownership to RQIA

- quality assurance measures to monitor and improve practice, as appropriate

- safeguarding children and adults at risk of harm

- responsibilities under health and safety legislation

- responsibilities under The Independent Health Care Regulations (Northern Ireland) 2005

- responsibilities under The Regulation and Improvement Authority (Independent Health Care) (Fees and Frequency of Inspections) (Amendment) Regulations (Northern Ireland) 2022

- responsibilities under the DoH Minimum Standards for Dental Care and Treatment (2011)

- responsibilities under The Ionising Radiations Regulations (Northern Ireland) 2017 and The Ionising Radiation (Medical Exposure) Regulations (Northern Ireland) 2018

- staff selection and recruitment procedures

- adherence to professional codes of conduct

- any court cases pending/disciplinary cases with employers/professional regulatory bodies

Registration of Mrs Shauna McHenry as the responsible individual and registered manager is approved.

6.0	Quality Improvement Plan/Areas for Improvement

This inspection resulted in no areas for improvement being identified.  Findings of the inspection were discussed with Mrs McHenry, applicant responsible individual, as part of the inspection process and can be found in the main body of the report.